## Solar tower
Solar tower 1 is located 1 m east of CP1 in a clearing
30 ft mast - two lengths extended to 21 feet
2 guywire sets at 10ft and 20 ft
Enclosure 1 is mounted at 1.5 m
Enclosure 2 is mounted at 2.0 m

Panels	Model	Mounting hardware	Wired to	Wired to	Lead length	Wattage	Notes
20 watt 1	Solarex…	complete brackets	Charge controller 1		4m(?)	20
20 watt 2	Solarex…	brackets found	Charge controller		4m(?)	20
20 watt 3	Solarex…	new brackets	Charge controller 2		10m(?)	30

Charge controllers	Model	Mounting hardware	Wired to	Wired to	Lead length
Controller 1	Morningstar SS-6-12V	velcro to interior of enclosure 1	Solarex20 1	Battery array
Controller 2	Morningstar SS-6-12V	""	Solarex20 2	Battery array

Batteries	Model	Mounting hardware	Wired to	Wired to	Lead length	AmpHrs	Notes
SunXtender 1	SunXtender…	Black battery box 1	Charge controller 2	Dataloggers (2)			Daisy chained together in box
SunXtender 2	SunXtender…	""	Charge controller 2	Dataloggers (2)			Daisy chained together in box
SunXtender sm1	SunXtender…	""	Charge controller 2	Dataloggers (2)			Daisy chained together in box
SunXtender sm2	SunXtender…	""	Charge controller 2	Dataloggers (2)			Daisy chained together in box
SunXtender sm3	SunXtender…	""	Charge controller 2	Dataloggers (2)			Daisy chained together in box

## Met tower
Met tower is located 16m east (and slightly downhill) from the solar tower
30 ft mast - two lengths extended to 21 feet
2 guywire sets at 10ft and 20 ft

Boom 1 (4ft length) is just above boom (~3.6m) and faces N/S
Boom 2 (6ft length) is at 3.5m, oriented east/west
Enclosure 1 is mounted at 1.5 m
Enclosure 2 is mounted at 1.4 m

Sensor	Model	ser num	Mounting hardware	Mounting location	Mounting ht/dir	Wired to	Lead length	Calibration 1	Calibration 2	Calibration 3	Calibration 4	Current draw	Notes
Temp/Humid	Vaisala HMP45AC	z1050105	RM Young beehive/mast	Mast, above enclosure	2m, directly above enclosure	Met1 CR23x	2m
Wind	Met One C034B	D2954	nurail on boom	Boom 1 - South end	3.7m, S end of boom 1	Met1 CR23x	3m (spliced & extended)
Precip	Texas Electronics TR525I	41503-609	hoseclamped/vert boom	Boom 1 - North end	3.7 m, N end of boom 2	Met1 CR23x	4.9m
PAR up	Li-Cor Quantum 190	Q27246	CM225	Boom 2 - East 0.9m	3.6m, straight up	Met1 CR23x		mult: -146.84	tcoff: .0036				Has 604 ohm millivolt adapter
PAR down	Li-Cor Quantum 190	Q33398	CM225	Boom 2 - East 0.9m	3.5m straight down	Met1 CR23x		mult: -158.23	tcoff: .0036				Has 560 ohm millivolt adapter or resistor
IR radiometer	Apogee SI-111	2363	CM230	Boom 2 - West 0.9m	3.5m - tilt 12d away from mast	Met1 CR23x		C2 m: 0.53303, b: 0.39440	C1 m: 76.35606, b: -9.88480	C0 m: 13419.73420, b: -72.61263			Calibration coefficients are saved in my drawer
Net radiation	REBS Q_7.1	Q03108		Boom 2 - East end	3.5+ m - boom juts to S	Met1 CR23x		pos: 8.65	neg: 10.54
Ultrasonic depth	Judd	2325	CM220	Boom 2 - West end	3.4m - (measure)	Met1 CR23x		output: 2.5v
Atmospheric pressure	Setra 278	3746248	screws in enclosure	Enclosure 1	1.5m	Met1 CR23x	35cm

Datalogger	Model	ser num	Mounting hardware	Notes
Met1	Campbell CR23x	5379 (inventory #276793)	Campbell enclosure on tower	See Met1 CR23x worksheet for wiring

Peripheral	Model	ser num	Mounting hardware	Notes
Relay	Campbell A6REL12	1615	Enclosure 1	Switches modem, Base, and Met radios

Comm device	Model	ser num	MAC address	Mounting hardware	Antenna	Notes
Wireless modem	Sierra wireless Raven XTG	910387742	3555633007164998 (IMEI)	Enclosure 2	Wilson antenna, boom 2
Base radio	Digi XbeePRO PKG		0013A2004063B85A	Enclosure 2	Whip antenna, boom 2
Met1 radio	Digi XbeePRO PKG		0013A2004063B8D6	Enclosure 1	Whip antenna, top of enc 1

## Forest Tower
The forest tower is 40.5 m east and downhill of the Met tower
The tower is a fencepost roughly 8 feet tall, driven 2 feet into the ground. It has one crossbar mounted at the top with a wind sensor at the end.
There are two IR radiometers mounted just below the crossbar and another on a post 5 meters east.
There is a datalogger at 4 feet in an enclosure with the Forest1 datalogger inside. This datalogger is wired to MUX1, MUX2, and MUX3

Boom 1 (4ft length) is just above boom (~3.6m) and faces E/W
Enclosure 1 is mounted at 1.5 m

Sensor	Model	ser num	Mounting hardware	Mounting location	Mounting ht/dir	Wired to	Lead length	Calibration 1	Calibration 2	Calibration 3	Calibration 4	Current draw	Notes
Temp/Humid	Vaisala HMP45A	W0530009	RM Young beehive/pole	Forest tower, above enclosure	1.6m	Forest1 CR23x
Wind	Met One C034B	C5805	nurail on boom	Boom 1	1.8m	Forest1 CR23x
IR radiometer	Apogee SI-111	2361	CM230	Post in treatment swath	Treatment edge	Forest1 CR23x		C2 m: 0.86384, b: 0.28300	C1 m: 62.23732, b: -4.82953	C0 m: 14338.44338, b: -12.73683
IR radiometer	Apogee SI-111	2362	CM230	Forest tower, above enclosure	Canopy	Forest1 CR23x		C2 m: 0.77405, b: 0.24962	C1 m: 68.07768, b: -2.25549	C0 m: 14222.34222, b: -72.48612
IR radiometer	Apogee SI-111	2364	CM230	Forest tower, above enclosure	Control edge	Forest1 CR23x		C2 m: 0.60645, b: 0.35482	C1 m: 73.83054, b: -9.16495	C0 m: 13076.13076, b: -82.13272

Datalogger	Model	ser num	Mounting hardware	Notes
Forest1	Campbell CR23x	6343 (inventory #279627)	Campbell enclosure on 8 ft post	See Forest1 CR23x worksheet for wiring

Peripheral	Model	ser num	Mounting hardware	Notes
Relay	Crydom D1D07		Enclosure 1	Switches forest radio
Multiplexer 1	Campbell AM16/32	2142	MUX1 Enclosure	See SM network sheet
Multiplexer 2	Campbell AM16/32B	6030	MUX2 Enclosure	See SM network sheet
Multiplexer 3	Campbell AM16/32	9600	MUX3 Enclosure	See SM network sheet

Comm device	Model	ser num	MAC address	Mounting hardware	Antenna	Notes
Forest1 radio	Digi XbeePRO PKG		0013A200406AA06E	Enclosure	Whip on top of enc

## Forest SM network
The forest soil network consists of three multiplexers connected to 10 soil moisture and temperature profiles
These multiplexers are controlled by the Forest1 datalogger and all soil sensors are read once per minute
MUX 1,2, and 3 are mounted in enclosures on shorter (5 feet?) fenceposts. These multiplexers connect to the soil profile sensors
Each multiplexer has two cables connecting it to the datalogger, a control cable (CTL) that provides power and switching signals to the multiplexer, and a communication cable (COM) that carries sensor signals and excitations

Multiplexer	Model	ser num	MUX Terminal	DL	DL Terminal	Line	Notes		Decagon wiring
Mux1	Campbell AM16/32	2142	RES	Forest1 CR23x	C1	CTL (Black)	Blue foil		Wire	Mux terminal
			CLK	Forest1 CR23x	C2	CTL (Green)	"		White (excitation)	ODD H
			GND	Forest1 CR23x	G (Power gnd 3)	CTL (Black)	Red foil, Jumped to term strip		Red (signal)	ODD L
			12V	Forest1 CR23x	12V (1)	CTL (Red)	Red foil, Jumped to term strip		Bare (Gnd)	COM GND

			COM ODD H	Forest1 CR23x	SW12V	COM (Red)	Red foil, 200ohm resistor in line
			COM ODD L	Forest1 CR23x	C5	COM (Black)	Red foil
			COM GND	Forest1 CR23x	signal gnd	COM (Green)	Blue foil
			COM EVEN H	Forest1 CR23x	SE 1	COM (Black)	Green foil		Campbell 107 wiring
			COM EVEN L	Forest1 CR23x	SE 2	COM (White)	"		Wire	Mux terminal
					Black(excitation)	Odd H
Mux2	Campbell AM16/32B	6030	RES	Forest1 CR23x	C1	CTL (Black)	Blue foil		Red (T signal)	Even H&L
			CLK	Forest1 CR23x	C3	CTL (Green)	"		Purple(signal gnd)	COM GND
			GND	Forest1 CR23x	G (Power gnd 3)	CTL (Black)	Red foil, Jumped to term strip		Clear(shield)	COM GND
			12V	Forest1 CR23x	12V (1)	CTL (Red)	Red foil, Jumped to term strip

			COM ODD H	Forest1 CR23x	SW12V	COM (Red)	Red foil, 200ohm resistor in line		Campbell 616/15 wiring
			COM ODD L	Forest1 CR23x	C6	COM (Black)	Red foil		Wire	Mux terminal
			COM GND	Forest1 CR23x	signal gnd	COM (Green)	Blue foil		Orange (enable)	ODD H
			COM EVEN H	Forest1 CR23x	SE 3	COM (Black)	Green foil		Red (12vdc)	Term 12v
			COM EVEN L	Forest1 CR23x	SE 4	COM (White)	"		Green (signal)	EVEN H&L
					Black (Ground)	COM GND
Mux3	Campbell AM16/32	9600	RES	Forest1 CR23x	C1	CTL (Black)	Blue foil		Clear (Shield)	Term G
			CLK	Forest1 CR23x	C4	CTL (Green)	"
			GND	Forest1 CR23x	G (Power gnd 4)	CTL (Black)	Red foil, Jumped to term strip
			12V	Forest1 CR23x	12V (2)	CTL (Red)	Red foil, Jumped to term strip		Sentek EnviroSCAN wiring
					Wire	Mux terminal
			COM ODD H	Forest1 CR23x	EX 3	COM (Red)	Red foil		Red(12vdc)	Term12v
			COM ODD L	Forest1 CR23x	C7	COM (Black)	"		Black(Ground)	TermG
			COM GND	Forest1 CR23x	signal ground	COM (Green)	Blue foil		Green (signal)	OddL (Control port)
			COM EVEN H	Forest1 CR23x	SE 5	COM (Black)	Green foil
			COM EVEN L	Forest1 CR23x	SE 6	COM (White)	"

Mux1 is located in an enclosure mounted on a fencepost 4 m NE of Profile #1d
Mux2 is located in an enclosure mounted on a fencepost 4 m SE of Profile #2d
Mux3 is located in an enclosure mounted on a fencepost 5 m W of Profile #7

	Sensor	Model	Mounting depth	Wired to	Channel	Lead length	Duff layer	Notes
Profile 1	Soil moisture	Campbell CS-615	5cm	Mux1	1	32m	10cm
18E 22N	Soil moisture	Campbell CS-615	20cm	Mux1	1	32m	10cm
	Soil moisture	Campbell CS-615	60cm	Mux1	2	32m	10cm

Profile 1d	Soil moisture/temp	Decagon EC-TM	5cm	Mux1	8	4.9m	12cm
17.5E 19N	Soil moisture/temp	Decagon EC-TM	20cm	Mux1	9	4.9m	12cm
	Soil moisture/temp	Decagon EC-TM	60cm	Mux1	10	4.9m	12cm

Profile 2	Soil moisture	Campbell CS-615	5cm	Mux2	1	32m	5cm
15E 64N	Soil moisture	Campbell CS-615	20cm	Mux2	1	32m	5cm	Bent but parallel
	Soil moisture	Campbell CS-616	60cm	Mux2	2	32m	5cm	Bent but parallel

Profile 2d	Soil moisture/temp	Decagon EC-TM	5cm	Mux2	8	4.9m	6cm
15E 66N	Soil moisture/temp	Decagon EC-TM	20cm	Mux2	9	4.9m	6cm
	Soil moisture/temp	Decagon EC-TM	60cm	Mux2	10	4.9m	6cm

Profile 3	Soil moisture	Campbell CS-616	5cm	Mux1	3	32m	4cm
31E 25N	Soil moisture	Campbell CS-616	20cm	Mux1	3	32m	4cm
	Soil moisture	Campbell CS-615	60cm	Mux1	2	32m	4cm	Out of order!

Profile 3d	Soil moisture/temp	Decagon EC-TM	5cm	Mux1	5	4.9m	2cm	via Junction box 2
30E 27N	Soil moisture/temp	Decagon EC-TM	20cm	Mux1	6	4.9m	2cm	via Junction box 2
	Soil moisture/temp	Decagon EC-TM	60cm	Mux1	7	4.9m	2cm	via Junction box 2

Profile 4	Soil moisture	Campbell CS-616	5cm	Mux2	2	32m	4cm	Bent but parallel
28E 54N	Soil moisture	Campbell CS-616	20cm	Mux2	3	32m	4cm
	Soil moisture	Campbell CS-616	60cm	Mux2	3	32m	4cm

Profile 4d	Soil moisture/temp	Decagon EC-TM	5cm	Mux2	5	4.9m	8cm	via Junction box 1
28E 50N	Soil moisture/temp	Decagon EC-TM	20cm	Mux2	6	4.9m	8cm	via Junction box 1
	Soil moisture/temp	Decagon EC-TM	60cm	Mux2	7	4.9m	8cm	via Junction box 1

Profile 5	Soil moisture/temp	Campbell CS-616 & 107	5cm	Mux3	1&7	25m, 8.25+	8cm	via Junction box 3
44E 13N	Soil moisture/temp	Campbell CS-616 & 107	20cm	Mux3	1&7	25m, 8.25+	8cm	via Junction box 3
	Soil moisture/temp	Campbell CS-616 & 107	60cm	Mux3	2&8	25m, 25m	8cm

Profile 6	Soil moisture	Campbell CS-616	5cm	Mux3	4	100ft
44E 55N?	Soil moisture	Campbell CS-616	20cm	Mux3	4	100ft
	Soil moisture	Campbell CS-616	60cm	Mux3	5	100ft

Profile 7	Soil moisture/temp	Campbell CS-616 & 107	5cm	Mux3	2&8	8.25m, 8.25m	4cm
55E 25 N	Soil moisture/temp	Campbell CS-616 & 107	20cm	Mux3	3&9	8.25m, 8.25m	4cm
	Soil moisture/temp	Campbell CS-616 & 107	60cm	Mux3	3&9	8.25m, 8.25m	4cm

Profile 8	Soil moisture	Campbell CS-616	5cm	Mux3	5	100ft
55E 55N?	Soil moisture	Campbell CS-616	20cm	Mux3	6	100ft
	Soil moisture	Campbell CS-616	60cm	Mux3	6	100ft

Profile 9	Soil moisture	Sentek EnviroSCAN	5cm	Mux3	11	100ft		via junction box 4
65 E 35N?	Soil moisture	Sentek EnviroSCAN	15cm	Mux3	11	100ft		via junction box 4
	Soil moisture	Sentek EnviroSCAN	40cm	Mux3	11	100ft		via junction box 4
	Soil moisture	Sentek EnviroSCAN	60cm	Mux3	11	100ft		via junction box 4

Profile 10	Soil moisture	Sentek EnviroSCAN	5cm	Mux3	12	100ft		via junction box 4
75 E 35N?	Soil moisture	Sentek EnviroSCAN	15cm	Mux3	12	100ft		via junction box 4
	Soil moisture	Sentek EnviroSCAN	40cm	Mux3	12	100ft		via junction box 4
	Soil moisture	Sentek EnviroSCAN	60cm	Mux3	12	100ft		via junction box 4

## Met1 CR23x
Datalogger	Model	ser num	Mounting hardware		Notes
Met 1	Campbell CR23x	5379 (inventory #276793)	Enclosure, base of Met tower

Sensor	Model	ser num	Mounting hardware	Mounting height/location	Wired to	Lead length	Calibration 1	Calibration 2	Calibration 3	Calibration 4	Current draw	Notes
Temp/Humid	Vaisala HMP45AC	z1050105	RM Young beehive/mast	2m	Met1 CR23x	2m
Wind	Met One C034B	D2954	nurail on boom	3.6m	Met1 CR23x	3.6m
Precip	Texas Electronics TR525I	41503-609	hoseclamped/vert boom	3.6m	Met1 CR23x	4.9m
PAR up	Li-Cor Quantum 190	Q27246	CM225	3.5m	Met1 CR23x		mult: -146.84	tcoff: .0036				Has 604 ohm millivolt adapter
PAR down	Li-Cor Quantum 190	Q33398	CM225	3.5m	Met1 CR23x		mult: -158.23	tcoff: .0036				Has 560 ohm resistor
IR radiometer - snowpack	Apogee SI-111	2363	CM230	3.5m	Met1 CR23x		C2 m: 0.53303, b: 0.39440	C1 m: 76.35606, b: -9.88480	C0 m: 13419.73420, b: -72.61263			Calibration coefficients are saved in my drawer
Net radiation	REBS Q_7.1	Q03108	custom	3.5m	Met1 CR23x		pos: 8.65	neg: 10.54
Ultrasonic depth	Judd	2325	CM220	3.5m	Met1 CR23x		output: 2.5v
Atmospheric pressure	Setra 278	3746248	velcro in enclosure 1	1.5m	Met1 CR23x	35cm

SE/DIFF terminals	Instrument	Wire	EX terminals	Instrument	Wire
Rank 1			gnd	Met One windset	Clear - Wind direction shield
SE1/1H	Vaisala T/H sensor	Yellow - T signal	EX1	Met One windset	Blue - Wind direction excit.
SE2/1L	Vaisala T/H sensor	Blue - H signal	EX2	IR radiometer (Met - 2363)	White - excitation
gnd	Vaisala T/H sensor	White, black, clear - sig ref, gnds	gnd
SE3/2H	Met One windset	Green - Wind direction sig	EX3
SE4/2L			EX4
gnd	Met One windset	White - Wind direction ref	gnd
SE5/3H	Setra Pressure	Blue - SE signal	CAO1
SE6/3L			CAO2
gnd	Setra Pressure	Yellow, Clear - sig ref, shield	gnd
SE7/4H	Judd u-s Depth	White - Diff High	P1	Met One windset	Red - Wind speed signal
SE8/4L	Judd u-s Depth	Brown - Diff Low	gnd	Met One windset	Black - Wind speed ref
gnd			P2	Texas Elec. Rain guage	Black - pulse signal
SE9/5H	PAR up	Green - pos	gnd	Texas Elec. Rain guage	White, Clear - ref, shield
SE10/5L	PAR up	Blue - neg	P3
gnd		Jumper from 5L	gnd
SE11/6H	PAR down	Bare - pos	P4
SE12/6L	PAR down	Clear - neg	gnd
gnd		Jumper from 6L

Rank 2			POWER OUT terminals
SE13/7H	Net radiation	Red - Diff High	GND	Ground to A6REL12 relay (2)	Black
SE14/7L	Net radiation	Black - Diff Low	5v	Power to A6REL12 relay (1)	Red
gnd	Net radiation	Clear - Shield, Jumper from 7L	GND
SE15/8H	IR radiometer (Met - 2363)	Green - signal	SW12	Vaisala T/H sensor
SE16/8L			GND	Setra Pressure	Black - power ground
gnd	IR radiometer (Met - 2363)	Blue - SE signal ref	12v	Power to A6REL12 relay (2)	Red - 12vdc supply
SE17/9H	IR radiometer (Met - 2363)	Red - Diff High	12v	Judd u-s Depth, Setra Pressure	Red - 12vdc supply
SE18/9L	""	Black - Diff Low	GND	Judd u-s Depth	Black, Clear - power ground, shield
gnd	""	Clear - shield	CONTROL I/O terminals
SE19/10H			C1	Setra Pressure	Green - enable
SE20/10L			C2	Judd u-s Depth	Green - enable
gnd			C3
SE21/11H			C4
SE22/11L			GND
gnd			C5	Control to A6REL12 relay	Green - enable
SE23/12H			C6
SE24/12L			C7
gnd			C8
			GND

## Forest1 CR23x
Datalogger	Model	ser num	Mounting hardware	Notes
Forest1	Campbell CR23x	6343 (inventory #279627)	Enclosure on Forest tower

Multiplexers
Mux1			Fencepost
Mux2			Fencepost
Mux3			Fencepost

Sensors	Model	Ser num	Mounting hardware	Mounting height/location	Wired to	Calibration 1	Calibration 2	Calibration 3	Notes
Temp/Humid	Vaisala HMP45A	W0530009	RM Young beehive/pole		Forest1 CR23x
Wind	Met One C034B	C5805	nurail on boom		Forest1 CR23x
IR radiometer	Apogee SI-111	2361	CM230	Treatment edge	Forest1 CR23x	C2 m: 0.86384, b: 0.28300	C1 m: 62.23732, b: -4.82953	C0 m: 14338.44338, b: -12.73683	Calibration coefficients are saved in my drawer
IR radiometer	Apogee SI-111	2362	CM230	Canopy	Forest1 CR23x	C2 m: 0.77405, b: 0.24962	C1 m: 68.07768, b: -2.25549	C0 m: 14222.34222, b: -72.48612
IR radiometer	Apogee SI-111	2364	CM230	Control edge	Forest1 CR23x	C2 m: 0.60645, b: 0.35482	C1 m: 73.83054, b: -9.16495	C0 m: 13076.13076, b: -82.13272

SE/DIFF terminals	Instrument	Wire	EX terminals	Instrument	Wire
Rank 1			gnd
SE1/1H	Mux1	COM Even H	EX1	IR radiometer (x3)	White - excitation
SE2/1L	Mux1	COM Even L	EX2	Met One windset	Blue - Wind direction excit.
gnd	Mux1 & Mux2	COM gnd	gnd	Met One windset	Clear - Wind direction shield
SE3/2H	Mux2	COM Even H	EX3	Mux3	COM Odd H
SE4/2L	Mux2	COM Even L	EX4
gnd	Mux3	COM gnd	gnd
SE5/3H	Mux3	COM Even H	CAO1
SE6/3L	Mux3	COM Even L	CAO2
gnd			gnd
SE7/4H	IR radiometer (Treatment - 2361)	Green - SE signal	P1	Met One windset	Red - Wind speed signal
SE8/4L	IR radiometer (Control - 2364)	Green - SE signal	gnd	Met One windset	Black - Wind speed ref
gnd	IR radiometers (2361 & 2364)	Blue - SE signal ref (2)	P2
SE9/5H	IR radiometer (Treatment - 2361)	Red - Diff High	gnd
SE10/5L	""	Black - Diff ref	P3
gnd	IR radiometer (2361)	Clear - shield	gnd
SE11/6H	IR radiometer (Control - 2364)	Red - Diff High	P4
SE12/6L	""	Black - Diff Low	gnd	Mux1, Mux2, Mux3	Shields
gnd	IR radiometer (2364)	Clear - shield

Rank 2	Instrument	Wire	POWER OUT terminals	Instrument	Wire
SE13/7H	IR radiometer (Canopy - 2362)	Green - SE signal	GND	Ground to Crydom relay	Black
SE14/7L			5v	Power to Crydom relay	Red
gnd	IR radiometer (Canopy - 2362)	Blue - SE signal ref	GND
SE15/8H	IR radiometer (Canopy - 2362)	Red - Diff High	SW12	Mux1, Mux2 & Vaisala T/H	COM Odd H & Red (enable)
SE16/8L	""	Black - Diff Low	GND	Mux1, 2	GND
gnd	""	Clear - shield	12v	Mux1, 2	12V
SE17/9H	Vaisala T/H sensor	White - T signal	12v	Mux3	12V
SE18/9L	Vaisala T/H sensor	Green H signal	GND	Mux3	GND
gnd	Vaisala T/H sensor	Blue, black, clear - sig ref, gnds	CONTROL I/O terminals
SE19/10H	Met One windset	Green - Wind direction sig	C1	Mux1, Mux2, & Mux3	RES
SE20/10L			C2	Mux1	CLK
gnd	Met One windset	White - Wind direction ref	C3	Mux2	CLK
SE21/11H			C4	Mux3	CLK
SE22/11L			GND
gnd			C5	Mux1	COM Odd L
SE23/12H			C6	Mux2	COM Odd L
SE24/12L			C7	Mux3	COM Odd L
gnd			C8	Control to Crydom relay	Green - enable
			GND